## Slide 2: Beam-on-Target Instrumentation: Tuning, Monitoring, and Beyond
Presented by Cyrille Thomas

## Slide 3: Outline
Beam on target properties
Instrumentation suite
Focus on IMG
Data collection and applications

## Slide 4: Beam Properties – ESS accelerator
Production:
Current: 62.5mA
Pulse: 2.86ms
Frequency: 14Hz
Energy: 2GeV
Peak power: 125MW
Average power: 5MW
Tunning:
Current: 1mA - 62.5mA
Pulse: 5µs - 2.86ms
Frequency: 1Hz - 14Hz
Energy: 216MeV - 2GeV
Peak power: 216kW - 125MW
Average power: 216mW - 5MW
Average beam size: 3mm

## Slide 5: Beam Properties on ESS Target
Beamlet current and position
Tuning:
Centre the beam: H ± 5mm, V ± 3mm
Set the lattice: beam size: H 13.5±1.4mm, V 5±0.5mm
Set the raster magnets: deflection amplitude H 60±6mm, V 20±2mm
Energy & current known (BPM & BCM)
courtesy from Y. Levinsen
2.86ms
62.5mA
pulse
62.5mA x 2GeV = 125MW

## Slide 6: High Power Proton Beam
Potential beam damage
Melting
Thermo-mechanical stress, fatigue
DT [ K / ms ] for materials
Material | Nominal beam | Static, nominal beamlet size | Static, half nominal beamlet size
SS | 5 | 56 | 222
Al | 3 | 34 | 134
W | 14 | 155 | 620
Mode | Nominal beam | Static, nominal beamlet size | Static, half nominal beamlet size
Slow Tuning (50µs) | 0.15 | 1.7 | 7
Production (2.86ms) | 8 | 94 | 375
DT [ K ] for Aluminium
Required:
Control beamlet size
Control raster beam
Interlock within one pulse period (71ms)
Risk of single pulse damage exist
Most condition leads to damage in many pulses
Example: 570MeV, 62mA
1.3cm x 0.5cm
0.65cm x 0.25cm
W to MW beam power
increasing pulse length, frequency and current
From tuning to production:

## Slide 7: Diagnostics Suites
  Complementary diagnostics: permits rapid and trustable verification
  Redundancy: permits high availability and reliability
  Fast and accurate online measurement supporting tuning, production and machine protection
GRID PBIP: 1D profile H and V
IMG TW: 2D profile
IMG PBW: 2D profile
APTM PBW: Halo, beam loss
APTM PBIP: Halo, beam loss

## Slide 8: Diagnostics complementarity and redundancy
 | IMG | GRD | APTM
Beam position |  |  | 
Beam size |  |  | 
Edges positions |  |  | 
Current Density Distribution |  |  | 
Beamlet position |  |  | 
Beamlet size |  |  | 
Beam loss and halo |  |  | 
Raster pattern |  |  | 
Online systems
Supports:
Beam tuning: set beam properties
Beam operation: detect errant beam condition
Direct observation
Indirect observation
All instruments in the suite play an important role, but let's focus on IMG

## Slide 9: T-IMG
System overview
lenses
TW
PBW
M1
M2
M4
M5
M6
M7
M3
beam
Source: luminescent material
PBW retro-reflector
Luminescent coating
PBW coating
TW coating
TW retro-reflector
optics
Camera
+
FPGA based image processing
real time analysis and interlock

## Slide 10
After:
alignment of the mirrors, camera, lighting system,
focussing right
set the TW on the reference position
0º
sector 1
sector 36
10º
beam
M1
RR
TW RR positions: reference
RR on moderator (missing)
dx = phase
dz
synchronisation: 10º / 14Hz period

## Slide 11
FPGA processing
Beam image
Fid. image
Aberration correction
Pre-conditioning
Errant condition check
14Hz
Beam timing ref
28Hz
triggers
SW: FPGA control settings and displays

## Slide 12
2 turns, synchronous speed 28 frames/s
1 turns, 2 frames / s
Camera speed is 28 fps:
one frames shows the beam
the other shows the RR fiducials when synchronously illuminated
RR fid. on sectors 1, 15, 26

## Slide 13
System magnification
L =83.3mm ± 0.57
Magnification: 0.0114 ± 0.7%
L = 61.6mm – from TW diameter
error: 2%
Early checks!
correction for geometrical aberration needed
L = 60.439 ± 0.57
Relative position for each RR fiducials
Standard deviation position | Value
vert. top RR | 0.061 mm
hor. top RR | 0.173 mm

## Slide 14
Lab system
Full replica of TW or PBW design
Laser raster system develop and deployed
Verify:
RR fiducial performance and lightning
Alignment procedure
Focussing procedure
Full vertical integration and performance
Lens
Camera
M1
M2
M3
M5
PBW
LED
Lab setup for development
M4
X: 3Hz -> 10kHz -> 30kHz
Y: 3Hz -> 10kHz -> 32kHz
Laser -raster

## Slide 15
Fid. position compensation
Image moments:
  centre cx, cy
  rms sizes sx, sy
Check conditions for FBIS interlock
  sx x sy > Ath
  |cx| <cx,max & |cy| <cy,max
FPGA processing:
-> FBIS interlock: stop beam production.
Raster frequency: fx =0.5Hz, fx =0.48Hz
Raster amplitude: ax = 24m, ay = 29mm
Camera exposure: 20ms
Short pulse equivalent: 0.26µs
Note: strong vertical jitter, due to screen vibration
Raster frequency: fx =500Hz, fx =482Hz
Raster amplitude: ax = 24m, ay = 29mm
Camera exposure: 20ms
Pulse equivalent: 0.26ms
Expected beam: large, stable
Expected beam: small, oscillating driven by raster
System test and verification

## Slide 16: … and beyond
Collection of data to build a database:
per pulse
current (BCM: current monitor)
total charge (BCM : current monitor)
position (IMG)
size (IMG / GRD)
2D image profile (IMG)
Raster centre of mass (GRD)
Beamlet size (GRD / IMG post analysis)
Post irradiation analysis
  all components in the target area are prone to be irradiated and subject to PIE -> invaluable data fully supporting analysis
Incident analysis
  in case of damage of a component, knowing the irradiation history is critical
Machine Learning
  new tools available bringing new angle in operation of the machine, data for ML training
commensurate data input for modelling the neutron source

## Slide 17: Concluding Remarks
In summary, beam-on-target instrumentation is essential not only for beam tuning but for ensuring safe and reliable operation at high power.
The combination of complementary diagnostics such as IMG, GRID, and APTM provides the necessary redundancy and accuracy to monitor beam properties in real time.
In particular, IMG enables direct 2D observation of beam position, size, and stability, supporting both operation and machine protection.
Finally, the systematic collection of beam data opens new opportunities for post-irradiation analysis, incident investigation, and the development of machine learning tools for future optimization.

## Slide 18
Thanks to all collaborators in this project, from ESS and from collaborating institutes.
E. Adli, H. Gjersdal, K. Sjobak - and other collaborators from University of Oslo, Norway
K. Michel, N. Weisend, M. Hartl, H. Kocevar, R.Kuriakose, M. Mohammednezhad, K. Rosengren, S. Ghatnekar Nilsson, R. Baron, M. Dominiak-Sumila, F. Nilen, A. Takibayev, and other collaborators from European Spallation Source ERIC
S. Joshi, S. Bjorklund - University West, Sweden
P. Jatczak and other collaborators from Institute of Electronic Systems, Warsaw University of Technology

## Slide 19: Thank you for your attention!

## Slide 20
Complementary slides

## Slide 21: Beam properties statistics over the commissioning
IMG02
Period from April to June
Beam area:
Range: 1-200mm2
Some peaks: most prominent at 50mm2
Can be below 25mm2
1.8% of the total pulses are below the lowest limit
The interlock status is changed to NOK for all these events
Beam position:
Covers the left part of the screen
Mostly close to the centre
Note: position interlock is for BPMs, and BLMs
Stability run
Beam area
Beam position
Notes:
Cavity tuning can only be done with armless beam (ESS probe beam mode)
Once tuned, stable condition can be found.

## Slide 22: Beam modes
Beam Mode | Current (mA) | Pulse Length (µs) | Frequency (Hz) | Average Power (kW) @ 2GeV | IMG |  | GRID |  | APTM | 
 |  |  |  |  | Tuning | MPS | Tuning | MPS | Tuning | MPS
Probe(low-low-low) | 6 | 5 | 1 | 0.06 | ✔︎ | ✗ | ✗ | ✗︎ | ✔︎ | ✗︎
Fast commissioning (low-low-high) | 6 | 5 | 14 | 0.84 | ✔︎ | ✗ | ✗ | ✗︎ | ✔︎ | ✗︎
RF test (low-high-low) | 6 | 50 | 1 | 0.60 | ✔︎ | ✗ | ✗ | ✗︎ | ✔︎ | ✗︎
Stability test (low-high-high) | 6 | 50 | 14 | 8.40 | ✔︎ | ✗ | ✗ | ✗︎ | ✔︎ | ✗︎
Slow commissioning (high-low-low ) | 62.5 | 5 | 1 | 0.625 | ✔︎ | ✗ | ✔︎ | ✗︎ | ✔︎ | ✗︎
Fast tuning (high-low-high) | 62.5 | 5 | 14 | 8.75 | ✔︎ | ✗︎ | ✔︎ | ✗︎ | ✔︎ | ✗︎
Slow tuning (high-high-low) | 62.5 | 50 | 1 | 6.25 | ✔︎ | ✗ | ✔︎ | ✔︎ | ✔︎ | ✗︎
Long pulse verification | 62.5 | 2860 | 0.033 | 11.92 | ✔ | ✔ | ✔ | ✔ | ✔ | ✔
Production | 62.5 | 2860 | 14 | 5005 | ✔ | ✔ | ✔ | ✔ | ✔︎ | ✔
From Slow tuning to Production
Increase pulse length
Increase frequency
increase current
All Tuning modes
Commission instrumentation
Production:
Raster monitoring required for beam permit

## Slide 23: GRID
3x Bias harps
1 Vert. harp: reads hor. profile
1 Hor. harp : reads vert. profile
Wire-clamps:
clamp
electrical contact
maintain tension
PBIP GRID slice

## Slide 24: GRID
FE, signal and bias
Hor. wires for Vert. profile
wire clamps and support
Assembled and under testing
Full vertical integration

## Slide 25: GRID
Control and sampling electronic
mTCA crate, timing, power, control
10 AMC board DAMCFMC2ZUP1 (CAEN ELS)
20 FMC-PICO-1M4-C
Readout 78 signals from wires
All channels individually process with FPGA
All signals transferred to concentrator FPGA board for real time processing of beamlet size and position
Interlock: evaluation every 1µs, delay 3µs
Firmware implementation
Based on (matlab) model of the instrument
Verified showing a stable and robust design

## Slide 26: GRID Model
aperture limit
beamlet size
beam centre: motion detection
Wire current sampling, 1MS/s
All samples synchronized
Beam abort if:
Beamlet size small than threshold values
No displacement detected
High current value detected on aperture edges
Tunning support: beamlet size and raster amplitude
Fast interlock: ~10µs
Operation range: Current > 30 mA

## Slide 27: Beam Diagnostics
Missions:
Support beam tuning
Support machine protection
in numbers:
  Beam position: 100
  Beam current : 21
  Beam loss: 348
  Beam profiles : 37
  Beam emittance: 2
  Destination stations: Faraday cups / beam stops: 6
  Tunning Dump and Target diagnostics instruments: 10
Total instrumentation: ~520 systems instances
Beam shaping: pulse choppers

## Slide 28: APTM
Measure beam outside defined aperture
p
e
Metallic plate
Secondary electron emission
p
Material
Heat load

## Slide 29: APTM
PBIP APTM
PBW APTM
Fully verified, pre-aligned to specs (±0.5mm)
Installed in PBIP in monolith vessel
Pre-assembled, fully verified and pre-aligned to specs (±0.5mm)
Installed in the PBW assembly
Metrology Survey

## Slide 30: APTM
Control and sampling electronic
Control software layer:
Control
Signal processing
Interlock
Archive
Display
Machine Protection
RCC1466 + Piggyback RTM
FBIS
FPGA: real time signal processing
AMC board DAMCFMC2ZUP1 (CAEN ELS)
mTCA crate, timing, power, control
FMC-PICO-1M4-C
TCs:
Beckhoff ECAT modules: EL3314-0002
1µs sampling
3-6µs sustained condition to trigger FBIS
3ms sampling
1ms trigger signal

## Slide 31: APTM
Blade signal during operation
APTM Model
Blade signal vs. beam position
TCs ∂T after each pulse
TCs steady state vs. beam offset
Blades
TCs
Proportional to current
Interlock within 10µs
Proportional to current density
Interlock within 10ms

## Slide 32
DumpL APTM
First signal and milestone reached!
Dump
at ~9m
APTM
IMG
Conceptual and final design (2018-2021)
  BI, collaborations (Lund University, ESS Mechanical, WUT)
Prototypes (2018, 2019 and 2020)
  BI, testing at JPARC, ESS support (mechanical, logistics …)
Manufacturing, assembling, installation (2019-2023)
  BI, ESS support (mechanical, vacuum, linac group Target group, procurement, rigging, …
Control system (2023-2026), Firmware, FPGA, low level software, EPICS: BI, ICS
Commissioning (now): BI, ICS, BP, Ops
  Verify instrument performance
  Verify protection function
  Conditions
    controlled beam: current, pulse, energy, size, position!
1
2
3
4
5
6
7
8
9
10
11
12
APTM
FE
Electronics:
sampling
signal processing
FBIS interface
Control
BL component
Long-haul cables
signal